## Slide 1: Odorant Measurement and Injection Control
Tomek Marchlewski, Eric Schaumloffel
Ohio Gas Association
March 31, 2016

## Slide 2: Presentation
Introduction
  Why do we measure?
  Presentation and analysis principle of the energyMEDOR analyzer
Performance and results validation
Summary

## Slide 3: Why do We Analyze for Sulfur Compounds?
Security (Health and Safety)
  Odorant injection control
  Pipeline integrity
Quality of supply
Cost concerns

## Slide 4: EnergyMEDOR
The EnergyMEDOR is an online analyzer for continuous monitoring of sulfur compounds (H2S, mercaptans, THT and TS) present in matrices such as Natural Gas.

## Slide 5: Analysis Principal
The EnergyMEDOR uses an electrochemical cell that allows the specific detection of sulfur species.
The detection is achieved by a gas-liquid reaction without altering the sample (no conversion of the sample).
An internal permeation unit is used to validate the results during each analysis (DMS permeation tube). This enables automatic internal calibration that allows for auto validation of the results.

## Slide 6: Overview
Oven (inside): injection valve, sampling loop and chromatographic column
Pressure regulator to adjust carrier gas flow
Calibration system (permeation tube inside the oven)
Selection valve (solenoid valve)
Adjustment valve to set the sampling flow
Pressure regulator to adjust the flow of the calibration system
Electrochemical detector and reservoir

## Slide 7: Analytical Operation
  Sample is introduced into the system.
  Sample gas travels through the sampling loop.
  Carrier gas travels through the columns and into detector.
  The sulfur compounds are separated in the column according to their affinity to the column film.
  The sulfurs are then detected by the wet cell through a gas-liquid reaction. The individual compounds are qualitatively and quantitatively identified by elution time and detector response.

## Slide 8: Analysis

## Slide 9: Software
  Analytical control
  Storage of data (calibrations, results and alarms)
  Full traceability for quality and audit trail purposes
  Real-time results transmitted via standard transfer protocols
  Set up of process alarms (injection monitoring)

## Slide 10: Sulfur Compounds Analyzed
Compound | Chemical formula
Hydrogen Sulfide Methyl Mercaptan (MTM) Ethyl Mercaptan (ETM) Dimethyl Sulfide (DMS) (iso) 2-Propyl Mercaptan (IPM) tert Butyl Mercaptan (TBM) Methyl Ethyl Sulfide (MES) (n) 1-Propyl Mercaptan (NPM) (sec) 2-Butyl Mercaptan (SBM) Diethyl Sulfide (DES) TetraHydroThiophene (THT) | H2S CH3-SH CH3CH2-SH CH3-S-CH3 (CH3)2-CH-SH (CH3)3-C-SH CH3CH2-S-CH3 CH3CH2CH2-SH CH3CH2CH(CH3)-SH CH3CH2-S-CH2CH3 C4H8S
Note: Compounds can be grouped to provide the odorization levels
as well as the relative index of odorization.
Compound | Chemical formula
Hydrogen Sulfide Methyl Mercaptan (MTM) Ethyl Mercaptan (ETM) Dimethyl Sulfide (DMS) (iso) 2-Propyl Mercaptan (IPM) tert Butyl Mercaptan (TBM) Methyl Ethyl Sulfide (MES) (n) 1-Propyl Mercaptan (NPM) (sec) 2-Butyl Mercaptan (SBM) Diethyl Sulfide (DES) TetraHydroThiophene (THT) | H2S CH3-SH CH3CH2-SH CH3-S-CH3 (CH3)2-CH-SH (CH3)3-C-SH CH3CH2-S-CH3 CH3CH2CH2-SH CH3CH2CH(CH3)-SH CH3CH2-S-CH2CH3 C4H8S

## Slide 11: Sulfur Compounds Analyzed

## Slide 12: Sulfur Speciation

## Slide 13: 3D Chromatograms

## Slide 14: Trend
0.5
0
lb/MMcf

## Slide 15: Trend
0
1
2
ppm and lb/MMcf

## Slide 16: Trend
2
1
0
ppm and lb/MMcf

## Slide 17: Annual Service
Permeation tube replacement
Permeation oven O-ring replacement
Rotor Replacement
Rotor housing cleaning
Sampling loop replacement
Detector cleaning
Flow and temperature check
Calibration check

## Slide 18: Summary
Detector is highly sensitive to sulfur compounds at ppm or ppb level
Very clear separation of the compounds
Stability of the results and validation of the results with standard permeation tube
Rugged industrial analyzer that needs very low maintenance
Odorant injection can be closely monitored allowing optimization of injection rates to reduce costs